## Slide 1: Concepts of med. equipment safety and risk management
Module 279-17-C Regulations, Standards and Ethics
Unit C 17.1 Medical Equipment Safety
  Safety is relative
  Hazards and risk
  Process of risk management
  Risk classification schemes
  Risk classification determinants
  Risk/benefit ratio for medical equipment
  Examples of risk classification based on determinants
17.1.1 Define the concepts of medical equipment safety and risk management
dr. Chris R. Mol, BME, NORTEC, 2017
see also: MEDICAL DEVICE REGULATIONS
Global overview and guiding principles
WHO 2003
©

## Slide 2: What is safe? What is risk?
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
a car
a race car
a truck
a bus
a bicycle
an airplane
a screw driver
a hammer
scissors
tongs / pincers
an electric drill machine
lathe machine
injection needle
suction machine
blood pressure machine
anesthesia machine
X-ray system
surgical knife
Sort on safety / risk level:
©

## Slide 3: Relative Risks in everyday life (USA data)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Activity | What is the risk that a person dies in any given year from this activity ?
Smoking 10 cigarettes per day | 
Influenza | 
Natural Causes, 40 years old | 
Road Accident | 
Accident at Home | 
Accident at Work | 
Hit by Lightning | 

## Slide 4: Relative Risks in everyday life (USA data)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Activity | What is the risk that a person dies in any given year from this activity ?
Smoking 10 cigarettes per day | 1 in 200
Influenza | 1 in 500
Natural Causes, 40 years old | 1 in 850
Road Accident | 1 in 8,000
Accident at Home | 1 in 26,000
Accident at Work | 1 in 43,500
Hit by Lightning | 1 in 10,000,000

## Slide 5: What is risk ?
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Values (such as physical health, social status, emotional well being or financial wealth) can be gained or lost when taking risk resulting from a given action, activity and/or inaction, foreseen or unforeseen.
Risk is the potential of losing something of value
Risk perception is the subjective judgment people make about the severity and/or probability of a risk, and may vary person to person.
Any human endeavour carries some risk,
but some are much riskier than others….
Uncertainty is a potential, unpredictable, unmeasurable and uncontrollable outcome,
risk is a consequence of action taken in spite of uncertainty.
Risk is the intentional interaction with uncertainty

## Slide 6: Safety and the chance of dying
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Your odds of dying by (in the USA):
  Cardiovascular disease: 1 in 2
  Smoking (by/before age 35): 1 in 600
  Car trip, USA coast-to-coast: 1 in 14,000
  Bicycle accident: 1 in 88,000
  Tornado: 1 in 450,000
  Train, coast-to-coast: 1 in 1,000,000
  Lightning: 1 in 1.9 million
  Bee sting: 1 in 5.5 million
  U.S. commercial jet airline: 1 in 7 million
Sources: Natural History Museum of Los Angeles County, Massachusetts Institute of Technology, University of California at Berkeley

## Slide 7: Risk is OK, sort of ….
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
In general, we are all ok with many of the risks of likelihoods in the range 1 in 10,000 to 1 in 10 million.
When we stop to consider many of these risks,
we recognize that they are potentially lethal and indeed could affect people such as ourselves,
but we do not change our behaviour to avoid them.
It is as if we recognize that there are just too many possible risks that might kill each one of us in our daily lives that we just adopt common sense and carry on living our lives.
“Every man dies, not every man really lives” (Braveheart)

## Slide 8: Exposure to Medical Radiation is not Insignificant
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Radiation is a well-established risk factor for Cancer.
Medical exposure to radiation is caused by X-ray / CT / NM examinations.
The amount of medical exposure (probably) increases the incidence of cancer in the population.
The amount of medical exposure to radiation therefore should be minimized.

## Slide 9: But such examinations can still be very ‘worthwhile’ (versus benefits) for individual patients
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
The biggest risk is taking no risks !
Investigation | Effective Dose (mSv) | Equivalent number of Chest X-rays | Equivalent period of natural radiation
Radiography |  |  | 
Chest | 0.02 | 1 | 3 days
Skull | 0.1 | 5 | 2 weeks
Abdomen | 1.5 | 75 | 9 months
CT examination |  |  | 
Brain | 2.0 | 100 | 1 year
Abdomen | 8.0 | 400 | 4 years

## Slide 10: Minimizing risks in Healthcare
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
How do we minimize the (unnecessary) risk that patients and hospital personnel run?
©
By exploring the sources of all potential accidents (hazards)
and
adapting our behaviour and our actions to avoid these.
Such adapted behaviour is made obligatory for all workers in the healthcare field
by means of
rules and regulations.

## Slide 11: Hazards and Risk
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
in this context, a ‘risk’ is a measure of the combination of
  the hazard
  the likelihood of occurrence of the adverse event
  the severity or overall impact
A ‘hazard’ is a ‘danger’: a potential source for an adverse event; something that could go wrong
a punctured tire
once every four years
a severe accident
A higher risk means more regulation….
The magnitude of the risk indicates
how much a hazard must be taken into account in practical life
©

## Slide 12: Hazards
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
car
hammer
X-ray system
List as many as possible hazards associated with a:
©

## Slide 13: Hazards and Risk (examples)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
 | hazard | how likely is it to happen (*) | how bad is it when it happens (*)
Injection needle | unintended puncture | +++ | ++
 | infection if non-sterile | + | ++
Suction machine | over-suction | + | +
Blood pressure machine | mercury spill | + | +
Anaesthesia machine | wrong dose administered | + | ++++
X-ray machine | patient damage from X-ray | + | +++
ECG machine | electrical shock through grounding error | + | ++++
There are usually many hazards per device!
©
(*) examples only

## Slide 14: Safety is a Risk Management issue
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
Starting points to consider medical device safety:
  Absolute safety cannot be guaranteed
  It is a risk management issue
  It is closely aligned with device effectiveness / performance
  It must be considered throughout the life span of the device
  It requires shared responsibility among the stakeholders
©
A ‘trade-off’ example is blood transfusion. In recent years, to reduce the risk of transmissible infection in the blood supply, donors with only a small probability of infection have been excluded. The result has been a critical shortage of blood for other lifesaving purposes, with a broad impact on patient care.

## Slide 15: Risk/benefit ratio for medical equipment
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
Medical devices should be designed and manufactured in such a way that:
  when used under the conditions and for the purposes intended and,
  where applicable, by virtue of the technical knowledge, experience, education or training of the intended users,
they will not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons,
provided that any risks which may be associated with their use constitute
  acceptable risks when weighed against the benefits to the patient and
  are compatible with a high level of protection of health and safety.
The first requirement of the
“Essential principles of safety and performance of medical devices”
recommended by the GHTF (SG1-N020R5) states that:
©
and maintained!!

## Slide 16: Risk needs to be assessed
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
All devices carry a certain degree of risk and could cause problems in specific circumstances.
Many medical device problems cannot be detected until extensive market experience is gained. For example:
  an implantable device may fail in a manner that was not predictable at the time of implantation
  a failure may reflect conditions unique to certain patients
  a component failure can be unpredictable or random
  …..
This estimate is often referred to as: the Risk Assessment
The current approach to device safety is to estimate the potential of a device becoming a hazard that could result in safety problems and harm.
©

## Slide 17: Risk needs to be assessed
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
Risk Assessment
begins with
Risk Analysis to identify all possible hazards,
followed by
Risk Evaluation to estimate the risk of each hazard.
In general, risk assessment is based on experience, evidence, computation, or even guesswork.
Risk assessment is complex, as it can be influenced by personal perception and other factors
such as cultural background, economic conditions, and political climates.
Risk Analysis
=
Hazard Identification
Risk Evaluation
=
Sizing the Risk
Risk Assessment
=
Risk Analysis plus Risk Evaluation
©

## Slide 18: Medical Device classification based on risk assessment
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
The degree of regulation imposed on any device is proportional to its potential hazard (risk class)
©
lecture 2

## Slide 19: ‘Medical Devices’ includes more than ‘Medical Equipment’
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
“Medical device” means any instrument, apparatus, implement, machine, appliance, implant, in vitro reagent or calibrator, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings for one or more of the specific purposes of:
  diagnosis, prevention, monitoring, treatment or alleviation of disease
  diagnosis, monitoring, treatment, alleviation of or compensation for an injury
  investigation, replacement, modification, or support of the anatomy or of a physiological process
  supporting or sustaining life
  control of conception
  disinfection of medical devices
  providing information for medical purposes by means of in vitro examination of specimens derived from the human body
‘Medical Device’ includes most medical items that are not drugs.
Regulation is mainly done at Medical Device level….
©
e.g. a bed versus a hospital bed,
a knife versus a surgical knife,

## Slide 20: Patient Safety?
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
In the United States, the public and the medical specialty of anesthesia were shocked in April 1982 by a USA television program entitled The Deep Sleep. Presenting accounts of anesthetic accidents, the producers stated that, every year, 6,000 Americans die or suffer brain damage related to such accidents.
In 2004, the Canadian Adverse Events Study found that adverse events occurred in more than 7% of hospital admissions, and estimated that 9,000 to 24,000 Canadians die annually after an avoidable medical error.
These and other reports have led the World Health Organization to estimate that one in ten persons receiving health care will suffer preventable harm.
Hospital Errors are the Third Leading Cause of Death in U.S.

## Slide 21: Patient Safety?
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
A Report on Quality and Safety in 2007 found that inadequate communication between healthcare providers, or between providers and the patient and family members, was the root cause of over half the serious adverse events in accredited hospitals. Handwritten reports, manual order entry, and poor legibility lead to substantial errors and injuries.
Other leading causes included inadequate assessment of the patient's condition, poor leadership or training.
Similarities and contrasts have been noted between the "cultures of safety" in medicine and aviation. Pilots and medical personnel both operate in complex environments, interact with technology, and are subject to fatigue, stress, danger, and loss of life and prestige as a consequence of error.
Given the excellent record of aviation in accident prevention, a similar medical adverse event system would include
    non-punitive reporting,
    teamwork training,
    feedback on performance and
    institutional commitment to data collection and analysis.

## Slide 22: Medical Device Risk Classification in the USA
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
The Food and Drug Administration (FDA) has established risk classifications for approximately 1,700 different generic types of devices.
Each of these generic types of devices is assigned to one of three regulatory classes based on the level of control necessary to assure the safety and effectiveness of the device. The risk classification is based on expert recommendation from members of 16 medical ‘panels’, together covering all devices.
Device classification depends on the intended use of the device and also upon indications for use.
For example, a scalpel's intended use is to cut tissue. A subset of intended use arises when a more specialized indication is added in the device's labelling such as, "for making incisions in the cornea".
Indications for use can be found in the device's labelling, but may also be conveyed orally during sale of the product.
Device classification is risk based, that is, the risk the device poses to the patient and/or the user is the major factor in the class it is assigned (see next page).
FDA Class I includes devices with the lowest risk and FDA Class III includes those with the greatest risk

## Slide 23: Risk classification determinants
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
In classifying devices, potential areas of hazard that are considered include:
  the degree of invasiveness
    an invasive device is usually considered to have higher potential hazard than an equivalent non-invasive device (e.g. there are invasive and non-invasive blood pressure monitors)
  duration of contact
    long versus short
  the body system affected
    more or less vital organs affected; e.g. the heart is considered vital
  local versus systemic effects
    systemic means involving the ‘total’ body
©

## Slide 24: Examples of risk classification (FDA)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Class I devices are deemed to be low risk and are therefore subject to the least regulatory controls.
For example, dental floss is classified as Class I device.
Class II devices are higher risk devices than Class I and require greater regulatory controls to provide reasonable assurance of the device’s safety and effectiveness. For example, condoms are classified as Class II devices.
Class III devices are generally the highest risk devices and are therefore subject to the highest level of regulatory control. Class III devices must typically be approved by FDA before they are marketed. For example, replacement heart valves are classified as Class III devices.

## Slide 25: Examples of risk classification (FDA)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
TITLE 21--FOOD AND DRUGS
CHAPTER I--FOOD AND DRUG ADMINISTRATION
DEPARTMENT OF HEALTH AND HUMAN SERVICES
SUBCHAPTER H--MEDICAL DEVICES
PART 870 -- CARDIOVASCULAR DEVICES
Subpart C--Cardiovascular Monitoring Devices
Sec. 870.2340 Electrocardiograph.
(a) Identification. An electrocardiograph is a device used to process the electrical signal transmitted through two or more electrocardiograph electrodes and to produce a visual display of the electrical signal produced by the heart.
(b) Classification. Class II (performance standards).
TITLE 21--FOOD AND DRUGS
CHAPTER I--FOOD AND DRUG ADMINISTRATION
DEPARTMENT OF HEALTH AND HUMAN SERVICES
SUBCHAPTER H--MEDICAL DEVICES
PART 870 -- CARDIOVASCULAR DEVICES
Subpart B--Cardiovascular Diagnostic Devices
Sec. 870.1130 Noninvasive blood pressure measurement system.
(a) Identification. A noninvasive blood pressure measurement system is a device that provides a signal from which systolic, diastolic, mean, or any combination of the three pressures can be derived through the use of tranducers placed on the surface of the body.
(b) Classification. Class II (performance standards).

## Slide 26: Examples of risk classification (FDA)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
Sec. 892.1000 Magnetic resonance diagnostic device.
(a) Identification. A magnetic resonance diagnostic device is intended for general diagnostic use to present images which reflect the spatial distribution and/or magnetic resonance spectra which reflect frequency and distribution of nuclei exhibiting nuclear magnetic resonance. Other physical parameters derived from the images and/or spectra may also be produced. The device includes hydrogen-1 (proton) imaging, sodium-23 imaging, hydrogen-1 spectroscopy, phosphorus-31 spectroscopy, and chemical shift imaging (preserving simultaneous frequency and spatial information).
(b) Classification. Class II.
Sec. 868.1100 Arterial blood sampling kit.
(a) Identification. An arterial blood sampling kit is a device, in kit form, used to obtain arterial blood samples from a patient for blood gas determinations. The kit may include a syringe, needle, cork, and heparin.
(b) Classification. Class I
Sec. 864.3600 Microscopes and accessories.
(a) Identification. Microscopes and accessories are optical instruments used to enlarge images of specimens, preparations, and cultures for medical purposes. Variations of microscopes and accessories (through a change in the light source) used for medical purposes include the following:
(b) Classification. Class I

## Slide 27: Risk Classification in the European Union (EU)
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
The classification is ‘rule based’: rules categorize medical devices according to their perceived potential hazards (see next page).
The classification of medical devices in the European Union is outlined in Annex IX of the Council Directive 93/42/EEC.
There are basically four classes, ranging from low risk to high risk.
  • Class I
  • Class IIa
  • Class IIb
  • Class III

## Slide 28
dr. Chris R. Mol, BME, NORTEC, 2017
Government Regulatory Agencies
©
EU Rule Based device classification (example)

## Slide 29: Examples of risk classification
Concepts of Safety and Risk Management
dr. Chris R. Mol, BME, NORTEC, 2017
©
(GHTF guidelines)

## Slide 30
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/