## Slide 1: Foetal (heart) monitor
Module 279 19 C Medical Instrumentation II
Unit C 18.2 Maintaining Gynaecology and Obstetrics equipment
18.2.3 Foetal Heart Monitor
dr. Chris R. Mol, BME, NORTEC, 2015
©
principles of operation
  function
  use
  scientific principles
construction
  components
  system diagram
  inputs/outputs
troubleshooting
  identifying common faults
  replacing components
  rectifying faults
safety considerations
  user and patient safety
  electrical safety

## Slide 2: Function
dr. Chris R. Mol, BME, NORTEC, 2015
Foetal Heart Monitor
©
With a fetal monitor (or: CardioTocoGraph, CTG) you can measure:
  the heart rate of the fetus with Doppler Ultrasound
  the strength and duration of uterine contractions of the mother during labour
Electronic fetal heart monitoring can be done during pregnancy, (from 26 weeks onwards), labor and delivery. Fetal heart rate monitoring is especially helpful for high-risk pregnancy conditions such as diabetes, high blood pressure, etc.
A normal fetal heart rate is between 110 to 160 beats a minute. The sound of the beat is usually strong and regular. It is normal to have some changes in the fetal heart rate during labor, but drastic changes in heart rate before or after a contraction may indicate that the fetus is in distress. A fetus with large changes in heart rate may need to be removed from the womb immediately, by Caesarean section.

## Slide 3: Fetal heart rate with ultrasound Doppler technique
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
For the Doppler technique, a nurse applies ultrasound gel to the end of a flat transducer and moves it across the abdomen until a good reflection is found from the fetal heart.
The Doppler technique is the most common as it is less prone to artefacts than the surface electrodes and it is easy to apply.
However, the transducer and accompanying electronics are expensive, may require repositioning during labor, and don’t resolve multiple fetal pregnancies well.
During pregnancy (from 12 weeks onwards) the heart rate of the fetus can be measured with a (stand alone) hand-held electronic Doppler ultrasound device.

## Slide 4: Uterine contractions
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
Uterine contractions are measured with a strain gage transducer or pressure sensor usually strapped with a belt around the abdomen positioned over the fundus (top) of the uterus. This indirectly records contractions through sensing changes in skin tension arising from the uterine muscular activity.
uterine
contractions
During labour, it is important to have some indication of how labour is progressing, and to monitor the build-up of contractions, both in terms of frequency and strength. Such contractions have a powerful impact on the fetus. Therefore, fetal heart rate is measured simultaneously.

## Slide 5: Use
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
During (uterine) contractions the heart rate will decrease and revert to previous levels after the end of the contraction. If there is a delay in the heart rate returning to its previous level it can indicate that there is fetal distress.
In labour, pregnant women are typically monitored continuously throughout labour, right up to delivery.
contractions
fetal heart rate

## Slide 6: System diagram
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor

## Slide 7: System diagram (example)
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
Find the Service Doc !

## Slide 8: PCB boards (example)
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
CPU Board
Front End Board

## Slide 9: Components
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor

## Slide 10: troubleshooting: User Error
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
User errors are common with fetal monitoring, e.g.
  Incorrect connections of the transducers
  Incorrect loading of the paper or loading of the wrong paper
  Lack of gel – gel can dry out over a period of time in long labour monitoring sessions; gels may flow away from where it’s needed.
  Loose belt – it is important to apply enough belt tension to ensure that firm contact with the skin is maintained.
  Transducer movement – it may have slipped due to maternal movement, pulling the cable, etc.
  Fetal movement – the fetus may have moved out of the ultrasound beam
  Poor transducer positioning. This is an acquired skill and should include palpation of the fetus to determine the best position for the transducer (over the fetal left shoulder blade).

## Slide 11: troubleshooting: Other common faults
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
The Doppler probe is the most sensitive part of both the handheld and bedside devices. When it breaks no output is heard, even when a stethoscope indicates the presence of a fetal heart beat. You can check the probe operation by gently tapping the probe surface about once per second. If this is not detected, there is certainly a probe problem.
Power supply problems, typically dead batteries, are common with the hand held fetal Doppler devices.
The probe consists of the transducer assembly, the cable, and a multi-pin fixable connector. The cable contains between 5 and 80 separate conductors. The most frequent malfunction occurs as a result of a break in one or more of the cable conductors. Such malfunctions are usually the result of mishandling of the cable or of soaking it with gel.
The probe is expensive and typically cannot be replaced. Fortunately, the cable can often be mended.
The face of the probe is usually an acoustic lens. It must be handled with care. Do not drop the probe, and avoid scratching the face with sharp objects. Keep the probe assembly clean of oil and gel. Always clean the probe and cable with a tissue or damp cloth, after finishing work.
Be sure that gel is being used between the ultrasonic transducer and the patient.

## Slide 12: Performance Check
dr. Chris R. Mol, BME, NORTEC, 2015
©
Foetal Heart Monitor
For the contraction monitor, stretch the belt, or very gently press on the transducer at a rate of about one gentle push every minute. Use a watch to verify the time between applications. The monitor should reflect your application pressure (approximately) and rate (accurately – about 10%).
The Doppler probe should work when placed on your chest, with the proper gel, near your heart. Check for the accuracy of both the digital display and the paper trace by comparing their output with a measure of your own heart rate from a watch. The two should be within 1 or 2 beats per minute of the correct rate.
If both contraction rate and fetal heart rate are reported accurately, then the device is ready to release to the floor.
electrical safety test !

## Slide 13
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/