## Slide 1: Insufflator
Module 279 19 C Medical Instrumentation II
Unit C 18.2 Maintaining Gynaecology and Obstetrics equipment
18.2.4 Maintain an insufflator
dr. Chris R. Mol, BME, NORTEC, 2015
©
principles of operation
  function
  use
  scientific principles
construction
  components
  system diagram
  inputs/outputs
safety considerations
  user and patient safety
  electrical safety

## Slide 2: Laparoscopy = keyhole surgery = minimally invasive surgery
dr. Chris R. Mol, BME, NORTEC, 2015
Insufflator
©
The procedure can be used to observe (diagnose) or to take tissue samples for biopsy through the tube. Also, a diseased organ can be removed with this approach.
The advantages of laparoscopy over traditional ‘open surgery’:
  a shorter hospital stay and faster recovery time
  less pain and bleeding after the operation
  reduced scarring
During laparoscopy, the surgeon makes one or more small incisions in the abdomen. These allow the surgeon to insert the laparoscope, small surgical tools, and a tube used to pump gas into the abdomen.
A laparoscope is a small tube that has a light source and a camera, which relays images of the inside of the abdomen or pelvis to a television monitor.
looking only

## Slide 3: Laparoscopy
dr. Chris R. Mol, BME, NORTEC, 2015
Insufflator
©
Laparoscopy is most commonly used in:
  gynaecology – the study and treatment of conditions affecting the female reproductive system
  gastroenterology – the study and treatment of conditions affecting the digestive system
  urology – the study and treatment of conditions affecting the urinary system
Laparoscopy is used to find problems such as cysts, adhesions, fibroids, and infection. Also e.g. sterilization, in vitro fertilisation (IVF) and appendectomy can be carried out.
The procedures are carried out under general anaesthesia.
‘observing, biopsy, surgery’

## Slide 4: Insufflator function
dr. Chris R. Mol, BME, NORTEC, 2015
Insufflator
©
An insufflator (Latin insufflatio "blowing on" or "into") is used to blow an inert gas into the peritoneal cavity to expand the workroom and reduce obstruction.
The insufflator is connected by a silicon or PVC tube of 2-4 meters length to the insufflator needle that is introduced in the abdomen.
For gas, CO2, N2O or room air can be used. CO2 is used most often because it can easily be absorbed in the blood. Usually 3 litres of gas is sufficient. It is introduced under low pressure.
During some procedures gas escapes from the abdomen. In that case, gas needs to be continuously kept insufflating to compensate for this. The insufflator allows setting of the gas pressure to maintain a certain pressure in the abdomen.
At the end of the clinical procedure, the gas is removed manually by applying pressure on the stomach wall and letting the gas escape through the abdominal incisions. Alternatively, it can be sucked back into the insufflator. The remaining gas is absorbed by the body.

## Slide 5: Insufflator use
dr. Chris R. Mol, BME, NORTEC, 2015
Insufflator
©
It is essential to protect the patient against too high gas pressure. Therefore, insufflators have a pressure relief mechanism and/or an audible high pressure alarm.
It must be avoided that body fluids from the patient flow back into the insufflator. To achieve this, the insufflator can be positioned higher than the patient.
Also it must be checked that the gas container does not get empty during the procedure and that the insufflator needle is not withdrawn from the patient before the insufflator is turned off.

## Slide 6: Use
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
pressure setting
flow setting
pressure gauge
flow gauge
gas volume used
outlet port

## Slide 7: Construction: physical view
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
Stryker 40L Insufflator
Use the gas heater and the heating tube to insufflate lukewarm gas (37° C).

## Slide 8
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
Stryker 40L Insufflator
Stryker 40L Insufflator
Construction: functional diagram

## Slide 9: Construction
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
Stryker 40L Insufflator

## Slide 10: Construction: Field Replaceable Units
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
Stryker 40L Insufflator

## Slide 11: Components
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator
gas connector
pressure regulation adjustment
Use a high-pressure tube to connect a CO2 gas cylinder to the rear gas inlet connection or connect to centralized CO2 gas supply.
Stryker 40L Insufflator

## Slide 12: Safety considerations
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator

## Slide 13: Safety considerations
dr. Chris R. Mol, BME, NORTEC, 2015
©
Insufflator

## Slide 14
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/