## Slide 1: Radiant Warmer
Module 279 19 C Medical Instrumentation II
Unit C18.3 Maintaining Paediatrics Equipment
principles of operation
  function
  use
  scientific principles
construction
  components
  system diagram
  inputs/outputs
troubleshooting
  identifying common faults
  replacing components
  rectifying faults
safety considerations
  user and patient safety
  electrical safety
18.3.3 Maintain a radiant warmer.
dr. Chris R. Mol, BME, NORTEC, 2015
©

## Slide 2: Radiant warmers
dr. Chris R. Mol, BME, NORTEC, 2015
Radiant Warmer
©
Radiant warmers are often used:
  to stabilize infants following delivery room resuscitation,
  for procedures which cannot be done in the confined space of an incubator,
  for transportation of the infant for surgical interventions.
Radiant warmers (and incubators) are used to maintain the body temperature of new-born infants. This is done to minimize the energy that is required for metabolic heat production. Hypothermia at birth is one of the most important risk factors for new-born mortality.

## Slide 3: Function
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
The heat output of radiant warmers is usually regulated by servo-control to keep the skin temperature constant at a site on the abdomen where a thermistor probe is attached.
Compared with incubators, the partition of body heat loss is quite different under radiant warmers. Radiant warmers increase convective and evaporative heat loss and insensible water loss but eliminate radiant heat loss or change it to net gain.
Its major disadvantage is the increase in insensible water loss produced by the radiant warmer. ‘Insensible' water loss (‘insensible’ because we are not aware of it) refers to water loss due to:
  water that passes through the skin and is lost by evaporation,
  water loss through breathing out
This is loss of pure water: there is no associated electrolyte loss.
The major advantage of the radiant warmer is the easy access it provides to critically-ill infants without disturbing the thermal environment.
temperature probe

## Slide 4: Manual mode
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
Manual mode is used for following situations:
  When you are anticipating a new baby to be brought under warmer care . Keep the warmer on with 100 % heater output, once the baby is arrived shift to servo mode.
  If baby is having fever, move to manual mode and make heater output minimum . If you keep a feverish baby irradiated in servo mode, the alarm will get activated time and again.
  When rapid warming of hypothermic baby has to be undertaken.
  In labour room when attending delivery, the manual mode alarms every 10 to 15 minutes indicating the bed is warm and ready. If the alarm is silenced, again it will reactivate after another 10-15 minutes.

## Slide 5: Heat and Temperature
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
All matter is made up of molecules and atoms. These atoms are always in different types of motion. The motion of atoms and molecules creates heat or thermal energy. All matter has this thermal energy. The more motion the atoms or molecules have the more heat or thermal energy they will have.
Some atoms and molecules move faster than others. Temperature is an average value of energy for all the atoms and molecules in a given system. Temperature is independent of how much matter there is in the system. It is simply an average of the energy in the system.
What is Temperature?
What is Heat?

## Slide 6: Heat Transfer
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
Heat transfer is the exchange of thermal energy between physical systems. There are 3 different mechanisms by which heat is transferred throughout our environment: convection, conduction and radiation
Convection is the transfer of heat from one place to another by the movement of fluids. Convection is usually the dominant form of heat transfer in liquids and gases.
(Thermal) Conduction is the transfer of internal energy by microscopic diffusion and collisions of particles (molecules, ..) within a body or between adjacent bodies. Conduction takes place in all materials (solids, liquids, gases,..). The heat spontaneously flows from a hotter to a colder body.
Radiation is a method of heat transfer that does not rely upon any contact between the heat source and the heated object. Heat can be transmitted though empty space by thermal radiation often called infrared radiation. This is a type electromagnetic radiation. No mass is exchanged and no medium is required in the process of radiation. Examples of radiation is the heat from the sun, or heat released from the filament of a light bulb.

## Slide 7: Components
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
Bassinet
Lamp (silicon quartz/infrared/ceramic/quartz crystal), to provide radiant heat, with parabolic mirror to spread heat
Skin probe, to attach to the baby’s skin, to measure skin temperature
Control panel (electronic/electrical/microprocessor based) with a collection of display and control features/knobs
Heater output control knobs to increasing or decreasing the heater output manually and heater output display
Temperature selection panel to select desired skin temperature and temperature display
Mode selector to selects manual or servo mode
Castor wheels
Alarms: air over-temperature/skin over-temperature/air sensor fail/power failure/ etc.
The radiant warmer consists of a bed/bassinet on which to place the infant, and an overhead heater that delivers radiant heat. A skin temperature probe monitors infant temperature. Heat output can be controlled manually or through baby mode (feedback mode) for thermo-regulation. Visual and audio alarms are present for safety.
neonatal skin temperature sensors
control panel

## Slide 8: System Diagram
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer

## Slide 9: Maintenance
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
Required frequency/time for preventive maintenance
  SI6 Months / 15 Minutes
  PM12 Months / 30 Minutes
  CAL12 Months / 60 Minutes
  SPR24 Months / 30 Minutes
  AI0 Months / 120 Minutes
Disinfection: when the equipment is in use, all external surfaces should be cleaned daily with an antiseptic solution. Spirit or other organic solvents must not be used to clean the glass side panels or display panel. For disinfection of the reusable probe, an isopropyl alcohol swab should be used.
Every seventh day, after shifting the baby to another cot, the used equipment should be cleaned thoroughly, first by light detergent solution and then by antiseptic solution. All detachable assemblies, are to be treated similarly.
Ongoing (preventive) maintenance is the key to increase the mean time between failures. The hospital biomedical engineer must regularly check equipment but the authorized company engineer must be called for preventive checks and major breakdowns. The control and power units should be calibrated every 4-6 months and thorough servicing should be done annually.
Temperature calibration should ensure sensitivity to ± 0.5° of the set value.

## Slide 10: Trouble shooting flow chart
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
service manual !

## Slide 11: Safety Considerations
dr. Chris R. Mol, BME, NORTEC, 2015
©
Radiant Warmer
Safety of patient: hyperthermia
Safety of personnel:
  Do not touch the protective grill under the radiant heater or the top of the heater assembly. These surfaces may be hot and a burn could result.
  Disconnect power to the Warmer and allow the heater rod to cool before cleaning to avoid the possibility of a burn.
  Disconnect the Warmer power cord and allow the unit to cool before replacing the alarm or observation lights.
Potential pitfalls of servo-controlled warmer:
In the event of probe getting displaced from baby’s abdominal skin, overheating of the baby will occur because the skin probe depicts air temperature and heater output keeps on increasing till probe temperature matches control temperature.
Also, repeated activation of alarm will occur when baby develops fever. In this situation, it is better to shift to manual mode with lowest heater output setting.

## Slide 12
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/