## Slide 1: X-ray System Function
Module 279 19 C Medical Instrumentation II
Unit C18.6 Maintaining Medical Imaging Equipment
Radiology fundamentals
X-ray tube
Production of x-rays
Nature and properties of x-rays
X-ray detection
Photographically correct x-ray images
Contrast
Sharpness
18.6.1 Explain how an x-ray machine generates and detects images
dr. Chris R. Mol, BME, NORTEC, 2016
©

## Slide 2: Nature and properties of x-rays
dr. Chris R. Mol, BME, NORTEC, 2016
©
X-rays are electromagnetic waves/quanta with more energy (and a shorter wavelength) than light. For this reason X-rays can penetrate into the (human) body and still get partly absorbed, making it a great wavelength for making images of the body.
Neighbouring Ultraviolet radiation does not go through the body, while Gamma rays go right through it with little absorption.
Radiology Fundamentals
X-rays are absorbed more by ‘dense’ materials, such as metals and bone; it will be little absorbed by gas (lungs). ‘Soft tissue’ absorption will be somewhat in the middle of this.
Medical X-ray uses X-rays with an energy of between 40 and 120 keV. The higher energies are used when more penetration is needed, i.e. in thicker objects. An X-ray filter (often made of Copper or Aluminum) is used to filter out the lower, non-productive X-ray energies.
Electro-Magnetic radiation Spectrum

## Slide 3: Production of x-rays
dr. Chris R. Mol, BME, NORTEC, 2016
©
X-rays are produced by ‘shooting’ high energy electrons on a piece of metal (tungsten or molybdenum), called the ‘anode’. The electrons are absorbed in this metal and cause the creation of X-ray radiation.
Radiology Fundamentals
You get high energy electrons by taking low energy electrons and accelerating these by means of a high electric voltage. This high voltage is created by the X-ray generator and applied between the anode (+), attracting the electrons, and the cathode (-).
Low energy electrons are created by creating an electric current through a special thin wire (tungsten filament). The purpose of the filament is to ‘boil off electrons’. This is similar to the generation of light by a light bulb when photons are sent out by its heated filament.
99% of the energy released is given off as heat and only 1% is in the form of X-ray radiation! Getting rid of the heat is a major issue in tube design and operation.
All this is accomplished in an X-ray tube.

## Slide 4: X-ray tube
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
X-ray Out
High Voltage In
X-ray Tube ‘Insert’; the inside is vacuum
X-ray Tube Container
protects the tube from mechanical shock and absorbs stray X-ray radiation
Anode disc rotation mechanism (‘rotor’)
anode disk
kathode
X-ray tubes wear out and need to be replaced regularly. Depending on use this will be done every 2-4 years or longer.

## Slide 5: X-ray tube
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
In high end X-ray tubes, the anode is rotating, so that it can absorb more heat. Temperatures of 3000oC can be reached.
Also, the anode may be cooled with water or with oil, that is circulated through the tube and cooled with a ‘heat exchanger’ outside the tube.
tubes with cooling fluid
The place on the anode where X-rays are generated is called the ‘focal spot’.

## Slide 6: X-ray High Voltage (HV) generator
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
The X-ray HV generator creates the high voltage required to accelerate the electrons travelling to the anode. The high voltage used is in the order of 80 kV = 80,000 Volts.
X-ray HV Generator
The X-ray HV generator also creates the (low voltage) electric current that is used to heat up the filament of the X-ray tube.
A special High Voltage cable is used to connect the X-ray generator and the X-ray tube.

## Slide 7: X-ray Control Unit
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
X-ray system
user interface
(or ‘generator desk’)
In simple X-ray systems, the voltage, filament current and X-ray pulse width are entered manually on a ‘generator desk’. Together, these parameters determine for a large part the resulting Image Quality.
In more modern and complex system, X-ray exposure control is performed automatically by the X-ray system. This is one reason that X-ray systems require that the user inputs what body part is being imaged. It uses this information to set the X-ray system parameters (settings).
Often, the control unit is located outside the X-ray room. If it is located inside the room, a radio-opaque (X-ray absorbing) protection screen, large enough to protect a standing operator, is part of the control unit. There is a lead glass window, so that the patient can be watched during the examination.
There must also be leaded protective aprons and gloves.

## Slide 8: X-ray detection
dr. Chris R. Mol, BME, NORTEC, 2016
©
That’s why an ‘intensifying screen’, made of fluorescent material, is mounted in front of the film. It converts the X-ray photons into visible light photons. The amount of light that it produces is proportional to the amount of incoming X-rays. The light photons are recorded by the photographic film.
Radiology Fundamentals
The same approach is used for Image Intensifiers and Flat Detectors: X-ray is transformed into Light before it is detected / recorded.
‘Photographic’ film is insensitive to X-ray: 98% of X-rays would travel through without interaction.

## Slide 9: X-ray detection: cassettes and films
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
Cassettes are light-proof, rigid containers that enclose the X-ray film to protect it from light. Within the cassette are two intensifying screens that fluoresce and produce visible light when irradiated by X-rays. The film is placed between the two intensifying screens, inside the cassette. The cassettes must be strong, rigid and durable. They must provide firm pressure so that there is good contact between the film and the screens, but must be easy to open in the dark.
Cassettes come in different sizes

## Slide 10: X-ray Image Quality
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
Most user complaints on a functioning X-ray systems are about reduced Image Quality. The factors that determine image quality are:
  spatial resolution or sharpness
  image contrast
  image noise
  image artefacts
Each of these factors can lead to reduced visibility of image details and to a sub-optimal/wrong diagnosis.
Some image quality reduction occurs in case of high X-ray absorption, as sometimes caused by obese patients.
This can lead to low contrast and high noise in the image. It can sometimes not be improved by improving the system.
Chest X-ray of normal (left) and obese (right) patient
Image artefact caused by wrong positioning of collimator in the top left of the image

## Slide 11: X-ray Image Quality
dr. Chris R. Mol, BME, NORTEC, 2016
©
Radiology Fundamentals
With a Image Quality Control program the image quality is checked and optimized regularly.
Some problems can be solved by a local technician, some require expert service support.
There is a long list of system factors that can cause reduced image quality. This includes:
  wrong setting of exposure parameters (wrong program selected)
  inexpert use of collimator, grid
  wrong film selection (wrong sensitivity)
  faulty chemical development bath (temperature, concentration, time)
  worn-out X-ray tube
  problems with Image Intensifiers (worn out, artefacts,…)
  etc.
Dental X-ray images taken with different X-ray parameter (kV, mA, s) settings
Exposure factors too high (too dark)
Exposure factors correct
Exposure factors too low (too light)

## Slide 12: Image Contrast
dr. Chris R. Mol, BME, NORTEC, 2016
©
Good contrast can show structures which are otherwise invisible
Radiology Fundamentals

## Slide 13: Image Noise
dr. Chris R. Mol, BME, NORTEC, 2016
©
low noise (left) and high noise (right) chest X-ray images
Radiology Fundamentals
X-ray images are often noisy when a (too) low X-ray dose has arrived at the detector.
This is comparable to making pictures with your smartphone in the evening darkness..

## Slide 14: Sharpness
dr. Chris R. Mol, BME, NORTEC, 2016
Radiology Fundamentals
©
Chest X-ray image with high (left) and low (right) sharpness
Sharpness/blurriness
depends on the size of the focal spot….

## Slide 15
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/