## Slide 1: Identify equipment used in Medical Imaging
Module 279 19 C Medical Instrumentation II
Unit C18.6 Maintaining Medical Imaging Equipment
Medical imaging department
Function
Personnel
Equipment in the department
Quality control in the department
Quality assurance in the department
18.6.1 Identify equipment in a medical imaging department
dr. Chris R. Mol, BME, NORTEC, 2016
©

## Slide 2: The ‘Medical Imaging’ or ‘Radiology’ Department
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Radiology is the medical specialty that applies Medical Imaging to support diagnosis and treatment of patients in the hospital. Radiologists are the ‘photographers’ or ‘image makers’ of the hospital.
Radiology is a separate department in the hospital since medical imaging is so complex that it requires a specialization (additional training) to create and interpret the images. There is a trend that more simple imaging techniques, such as ultrasound, are located at - and owned by - ‘clinical departments’.
In general, patients are not ‘treated’ by Radiologists. Patients are referred to the Radiology department by ‘clinical departments, such as Cardiology, Neurology, Gynaecology, etc.
Radiology is (mostly) a Diagnostic department. Images are collected to support the diagnosis of a patient.
For a full diagnosis, not only images are required but also other patient data, such as patient history and medication. This diagnosis is made by the ‘referring physician’.

## Slide 3: The ‘Medical Imaging’ or ‘Radiology’ Department
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
The Radiology department can contain imaging techniques such as:
  X-ray radiography
  x-ray fluoroscopy
  ultrasound
  computed tomography (CT)
  nuclear medicine (NM)
  magnetic resonance imaging (MRI)
In advanced healthcare systems, ‘imaging’ is combined with therapeutic ‘interventions’ to create an ‘Interventional Radiology’ department, executing ‘image guided treatment’ procedures. Interventional image-guided treatments include biopsy, drainage, embolization, ablation, and others.

## Slide 4: Main activities in the department
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Scheduling the patient for the appropriate test
Provide proper equipment to perform the test (including stable reliable electricity, etc.)
Appropriately trained personnel to operate the equipment
System to visualize, store and transmit the imaging data
Radiologist to interpret imaging data
Generation of a report of the findings from the imaging data
System to transmit the report to the referring healthcare provider
Equipment maintenance program
System of continuing education and quality assurance to make sure all equipment is properly maintained and personnel are appropriately trained
Financial method of reimbursement to make the system viable and sustainable
Radiology in advanced healthcare systems has become mainly digital with all interpretations occurring from a computer workstation enabling advanced manipulation of images. This requires new staff, responsibilities, and skill sets.

## Slide 5: Personnel in Medical Imaging Department
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Personnel at the Medical Imaging Department can include:
  Administrators: to schedule patients, record payments, manage human resources, ….
  Nurses: Perform procedures on patients that require imaging exam, ensure patient safety,
  Radiographers or Radiologic Technologists: operate imaging hardware, transmit images (print or electronic) to radiologists
  Physicists: manage quality (resolution, calibrations) of imaging equipment, manage safety (radiation dose
  Radiologists: are Medical Doctors. Interpret/read the images. His/her report is transmitted to the clinician who requested the imaging.
Radiographer
Radiologist

## Slide 6: Imaging Modalities
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
In the following slides we will go over the main imaging modalities in the Medical Imaging department.
In subsequent lectures we will go into further detail on X-ray an Ultrasound images
A final lecture in this ‘Medical Imaging’ section will elaborate some more on
‘advanced’ 3D imaging techniques, including CT, MR, Nuclear Medicine and Tele-radiology.

## Slide 7: Radiography
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Diagnosis ?
Radiography is the oldest medical imaging modality since end 19th century. Its invention brought the start of the radiology department and was an important factor in the creation of the modern hospital.
Plain radiography was the only imaging modality available during the first 50 years of radiology. Due to its availability, speed, and lower costs compared to most other modalities, radiography is today often the first test of choice in radiologic diagnosis.

## Slide 8: Traditional film-screen radiography
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
an X-ray tube generates a beam of X-rays, which is aimed at the patient with the help of a collimator (shutters).
Some rays pass through the patient, while some are absorbed or scattered.
The X-rays that pass through the patient are filtered through a device called a grid or X-ray filter, to reduce scatter radiation by absorbing X-rays that do not come straight from the X-ray tube.
‘Automatic exposure control’ (AEC) uses X-ray detectors to switch off the X-ray tube when sufficient X-rays have reached the detector.
X-rays then strike an undeveloped film, which is held tightly to a screen of light-emitting phosphors in a light-tight cassette.
The film is then developed chemically so that the image on the film becomes visible.
The dark parts on the film occur where a high X-ray dose has arrived. This indicates patient tissue with a low absorption, such as the lungs. The light parts indicate tissues of high absorption, such as bones or metal.
X-ray films come in different sizes, dependent on the imaging task

## Slide 9: Digital Radiography
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
In digital radiography, X-rays strike a semi-conductor based Flat Detector that converts the X-ray intensity into digital information, which is transmitted and converted immediately (no need for chemical processing!) into an image displayed on a computer screen.
Over the last 10-20 years, film-screen radiography is being replaced by digital radiography. Just as film-based consumer photography has been replaced by digital photography.
a detachable Flat Detector can easily be repositioned
Digital Chest X-ray Imaging

## Slide 10: X-ray Fluoroscopy
X-ray ‘C-arm’ - with Image Intensifier - for use during surgery
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Fluoroscopy uses X-ray to make film, rather than still photo’s, to investigate objects that move. This is called dynamic imaging.
Conventionally, this is achieved by receiving the X-ray radiation on a fluorescent screen, coupled to an image intensifier tube. The optical light output of the Image Intensifier is picked up by a CCD camera and displayed on a monitor.
However, also for dynamic imaging, detector technology has moved to Flat Detector technology.
Diagram of Image Intensifier based system for dynamic X-ray imaging
Image Intensifiers with different sizes

## Slide 11: X-ray Fluoroscopy: some applications
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
A common application of fluoroscopy is the investigation of the stomach and intestines. For this, radiocontrast agents (e.g. Barium Sulphate) are swallowed by the patient to outline the anatomy and functioning of the organs. These agents strongly absorb X-rays and show up in the image as white areas.
swallowing contrast material
intestines
aorta and kidneys
In another application, Iodine-based contrast material is injected to delineate the blood vessels. These are investigated for potential blockages and other anomalies.
‘Digital Image subtraction’ and other digital image processing techniques are used to improve visibility of the interesting details. The contrast is then usually inverted so that the contrast medium shows up dark.
blood vessels in brain

## Slide 12: Over-projection: the 3D problem in Radiography
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Lateral Image:
In the nose !
Frontal Image:
Where is the absorbing object located ?

## Slide 13: Ultrasound
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Medical ultrasonography uses ultrasound to visualize soft tissue structures in the body in real time.
The quality of the images obtained using ultrasound is highly dependent on the skill of the person (ultrasonographer) performing the exam and the patient's body size (obesity).
Advantages of ultrasound imaging are that it provides images in real-time, it is portable and can be brought to the bedside. It is substantially lower in cost, and it does not use harmful ionizing radiation. This is a good reason why it is used so much in gynaecology/obstetrics to diagnose and measure progress of the fetal growth.
Examinations of larger, overweight patients may have a decrease in image quality as their subcutaneous fat absorbs more of the sound waves. This results in fewer sound waves penetrating to underlying organs and reflecting back to the transducer, resulting in poorer quality image. Ultrasound is also limited by its inability to image through air pockets (e.g. lungs) or bone.

## Slide 14: Ultrasound applications
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
Ultrasound is routinely applied in many parts of the body and for many applications, including
gynaecology and neonatology
cardiology, to detect ventricular size and heart valve anomalies
neurology for blood flow assessment
nephrology (kidney, bladder)
Ultrasound systems can deliver static or dynamic (moving) images (e.g. heart valves).
bladder with
cancer
kidney with cancer
blood vessels
heart valves
liver with cirrhosis
(too much alcohol)

## Slide 15: X-ray computed tomography
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
CT imaging uses X-rays in conjunction with computing algorithms to image the body. In CT, an X-ray tube opposite an X-ray detector (or detectors) in a ring-shaped apparatus rotate around a patient, producing a computer-generated cross-sectional image (tomogram). Tissues are not super-imposed, as in conventional Radiography. CT offers improved resolution, at the cost of a higher absorbed radiation dose.
The introduction of computed tomography in the early 1970s revolutionized diagnostic radiology by providing clinicians with images of real three-dimensional anatomic structures.
Continuing improvements in CT technology, including faster scanning times and improved resolution, have dramatically increased the accuracy and usefulness of CT scanning, which may partially account for increased use in medical diagnosis.
CT scanner
CT image cross section of patient head

## Slide 16: Quality Assurance and Quality Control
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
The Quality Assurance (QA) Program is a program designed by management to assure quality of a product or service. Such a program can have wide-ranging aspects, including customer feedback, employee empowerment, and quality control.
It is essential that one person at a given facility, the QA Coordinator, is in charge of maintaining the QA program and that (s)he is allotted the time, equipment, and space necessary to carry out the duties that are required. The facility’s QA Coordinator may choose to assign specific duties to other personnel but should maintain oversight and realize that consistency is compromised when several people share the responsibility of carrying out these tasks.
In Quality Assurance you aim to avoid defects. It is about prevention. It is a process based approach: processes are described and analysed so that problems cannot occur.
In Quality Control, you try to find defects and correct them. It is about detection of problems. It is used to verify the quality of a device.

## Slide 17: Quality control
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
A Quality Control Program allows a facility to efficiently monitor the components of the imaging process. It consists of a series of standardized tests developed to detect changes in X-ray equipment function from its original level of performance. The results of such tests, when carried out routinely, allow prompt corrective action to maintain image quality.
As an example, QC program for radiography equipment should include regular checks of:
  the film processor unit
  the darkroom
  the view boxes
  the film and chemical storage
  the intensifying screen cleaning
  the screen-film contact test
  the integrity of lead aprons and gloves
These actions typically involve measurement of a process variable, checking the measured value against a limit, and performing corrective action if the limit is exceeded.

## Slide 18: Quality control
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
A good QC program at the radiology department executes such a program for each medical imaging device…..
For X-ray systems, quality control is particularly important since radiation exposure cannot be seen and can be easily neglected without the presence of qualified personnel who are actively monitoring risk.

## Slide 19: QC Programs are not totally free…. (example only)
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
In order to perform the QC tests outlined in this manual, an initial outlay of about $1,600 will be necessary. Minimal annual costs can be expected to keep your QA and QC programs running.
  Sensitometer, approximately $700,
  Densitometer, approximately $700,
  Box of film (clinically used),
  Aluminum step wedge, approximately $70,
  Brass or copper mesh screens large enough to cover largest cassette in use at facility, approximately $10,
  Measuring tape,
  Non-mercury thermometer, approximately $10,
  Cleaning equipment for screens, cassettes and darkroom,
  Fluoroscopic test tool, approximately $100
Once you have become proficient at performing the tests the time spent will be minimal.
  Daily tests should take about 5 minutes to perform and should be done prior to the first patient of the day.
  Monthly tests will add an additional 10 minutes to the daily tests.
  Quarterly tests will take about 45 minutes to perform.
  The semi-annual test for darkroom fog should take no more than 5 minutes to perform and analyze.
  The annual tests will probably take 1 to 2 hours to perform.

## Slide 20: Quality Control Manual
dr. Chris R. Mol, BME, NORTEC, 2016
Medical Imaging
©
A QC Manual should be created and reviewed at least annually. The manual should include the facility’s objectives, QC instructions, QC results, and personnel responsibility.
Items that should be included in a QC Manual are:
  A list of the tests to be performed and the frequency for each test, including acceptable test limits, test procedures, maintenance, and service records.
  A list of equipment to be used for testing. Policy and procedures for QC tests as well as for the facility.
  Sample forms.
see for further detail:
QUALITY CONTROL RECOMMENDATIONS FOR DIAGNOSTIC RADIOGRAPHY, July 2001 available at http://www.crcpd.org/Pubs/QC-Docs/QC-Vol3-Web.pdf
Questions asked during a review might include:
  is image quality maintained at the desired level?
  Do all personnel meet required or established qualifications? Does any equipment need to be replaced?
  Do any QC procedures need to be changed or updated?
  Are personnel adequately performing assigned tasks?
  Is patient and personnel radiation exposure as low as reasonably achievable?
  Are all documents up-to-date and accurate?

## Slide 21
END
The creation of this presentation was supported by a grant from THET:
see https://www.thet.org/